## Slide 1: Automated platform for rapid multifunctional materials characterization under dynamic programmable environments
Eric Muckley, Ilia Ivanov
Abstract
Novel materials must be screened for stability and functionality under a broad range of ambient environmental conditions before they can be deployed in real-world applications. Multidimensionality of the environmental parameter space (temperature, pressure, humidity, gas/vapor composition) and common characterization modes (electrical, optical, gravimetric, viscoelastic) creates unique challenges for efficient data processing and experimental design. We developed an integrated computer-controlled system which enables simultaneous multi-sample screening of electrical, optical, gravimetric and viscoelastic properties of materials under programmable dynamic environments. Custom software performs on-the-fly data analysis and modeling, real-time iterative feedback for dynamic experimental settings, and rapid visualization of experimental results. The system reduces the time needed for researcher-directed experiment planning and data processing from hours to minutes, enabling high-throughput characterization of functionality and environmental response of a broad range of materials including oxides, nanotubes, two-dimensional materials, biomolecules, perovskites, polymers and ionic liquids. Here we demonstrate capabilities of the system, describe operational algorithms, and provide examples of high-quality multifunctional data acquired with minimal human intervention.

## Slide 2
Multifunctional characterization
High-throughput synthesis
Simulation and modeling
Materials database
Theory
Experiment
(a)
(b)
(c)
In (a) color scheme need to have multifunctional characterization standing off the rest since we are talking about it in (b) and (c)
Multifunctional characterization in traditional workflow may need to be duplicated with offset to show that the workflow is only for a single functionality. To test other functionality, one must to replicate the same workflow but for different instrument
Computer-controlled multifunctional response workflow removes limitation of transitional workflow by giving full control of instruments, experimental conditions and parameters through real-time iterative feedback
Traditional workflow
Automated workflow
Data acquisition
Human processes and models data
Human interprets and presents data
Manual experiment planning
Human adjusts experimental settings
Data acquisition
Computer processes and models data
Computer presents data
Real-time iterative feedback
Computer adjusts experimental settings
Each material functionality

## Slide 3
(a)
(b)
In (a) color scheme need to have multifunctional characterization standing off the rest since we are talking about it in (b) and (c)
Multifunctional characterization in traditional workflow may need to be duplicated with offset to show that the workflow is only for a single functionality. To test other functionality, one must to replicate the same workflow but for different instrument
Computer-controlled multifunctional response workflow removes limitation of transitional workflow by giving full control of instruments, experimental conditions and parameters through real-time iterative feedback
Traditional workflow
Automated workflow
Data acquisition
Human processes and models data
Human interprets and presents data
Manual experiment planning
Human adjusts experimental settings
Data acquisition
Computer processes and models data
Computer presents data
Real-time iterative feedback
Computer adjusts experimental settings
Each material functionality

## Slide 4
Optical: THz
Electrical: 0 – 10 MHz
Gravimetric / viscoelastic: 5 – 100 MHz
steady-state, ionic, dielectric
optical constants, vibrational, electronic
sorption kinetics, swelling
(a)
(c)
(b)
Environmental control
Source-measure unit
Impedance analyzer
QCM controller
Optical spectrometer
Sample
Sample
Sample
Vacuum chamber
Sample
Feedback
Control software
(d)

## Slide 5
Initialize experiment
Data analytics and modeling
Data acquisition
Real-time feedback
Design of experiment
Measure QCM conductance spectra
Set frequency range, bandwidth, crystal harmonics
Deconvolute and fit peaks, calculate Δf, ΔD, viscoelastic properties
Measure spectra
Set wavelength range, integration time
Calculate SNR, deconvolute and fit peaks
Measure I-V, C-V, impedance
Set bias range, step size, frequency range
Calculate conductivity, capacitance
Adjust sweep rate, bias, step size, frequency range
Adjust wavelength range, integration time, fit parameters
Set pressure, temperature, RH, gas/vapor flow
Gravimetric/viscoelastic
Optical
Electrical
Environmental
Adjust frequency range, bandwidth, fit parameters
Select next condition based on material response or user-defined sequence
False
Material response has stabilized
True
Maintain constant conditions
Update model of multifunctional material response to environment (statistics, physical models, autoML regression)
Export data, export models, create publication-quality plots
An extra blocks with the start of experiment are required, followed by selection of preset condition of user-defined for selected set of functionalities…
Start select functionalities/data streams select instruments - program environmental steps set instruments (default) or user-defined.
User defined are shown in the flow chart
What’s data analytics behind “Material stabilized” decision, probably better signal stabilization?
Can it be done for all channels together for one channel or for selected data channels?
Please add arrow pointing the output of each data channel.
Currently they cycled within independent channel. At which point the software makes a decision  pass to modeling “update model” ?. Include autoML which was used in text, also python scripts.
Block ” Adjust sweep rate bias step size” is for I-V and C-V , but not for impedance.
Select user-defined or automatic sequence
Select material functionalities/experiments
Select instruments
Start software, choose directory for storing results

## Slide 6
(a)
(b)
(c)
(d)
(e)
(f)
Added real axes titles and units

## Slide 7
(a)
(b)
(c)
(d)

## Slide 8
(b)
(c)
200 mV/s
50 mV/s
25 mV/s
(a)
2%
95%

## Slide 10
(a)
(b)
(c)
(d)
(e)

## Slide 12: UNUSED / SUPPLEMENTAL FIGURES

## Slide 13
resonance
spur
spur
(a)
(b)
(d)
(c)

## Slide 15
(a)
(b)
(c)
(d)

## Slide 16
2
25
50
75
95
RH (%)
200 mV/s
50 mV/s
25 mV/s
(a)
(b)
(c)

## Slide 17
AC electrical
DC electrical
Gravimetric/ viscoelastic
Optical
real dielectric constant
imaginary dielectric constant
current @ +1V bias
freq. shift @ 5th overtone
shear modulus @ 5th overtone
index of refraction @ 820 nm
extinction coef. @ 820 nm

## Slide 18
real dielectric constant
relative humidity
imaginary dielectric constant
current @ +1V bias
freq. shift @ 5th overtone
shear modulus @ 5th overtone
index of refraction @ 820 nm
AC electrical
DC electrical
Gravimetric/ viscoelastic
Optical
extinction coef. @ 820 nm

## Slide 19
Data analytics and modeling
Data acquisition
Real-time feedback
Design of experiment
Measure QCM conductance spectra
Set frequency range, bandwidth, crystal harmonics
Deconvolute and fit peaks, calculate Δf, ΔD, viscoelastic properties
Measure spectra
Set wavelength range, integration time
Calculate SNR, deconvolute and fit peaks
Measure I-V, C-V, impedance
Set bias range, step size, frequency range
Calculate conductivity, capacitance
Adjust sweep rate, bias, step size
Adjust wavelength range, integration time, fit parameters
Set pressure, temperature, RH, gas/vapor flow
Gravimetric/viscoelastic
Optical
Electrical
Environmental
Adjust frequency range, bandwidth, fit parameters
Select next condition based on material response
False
Material stabilized
True
Maintain constant conditions
Update model of multifunctional material response to environment
Export model, publication-quality plots

## Slide 20
(a)
(b)
(c)
(d)
(e)
(f)

## Slide 21
Optical
Electrical
Hz
THz
Gravimetric / viscoelastic
steady-state / ionic
dielectric
vibrational
optical constants
electronic
sorption kinetics
swelling
(a)
(c)
MHz
(b)
Environmental control
Source-measure unit
Impedance analyzer
QCM controller
Optical spectrometer
Sample
Sample
Sample
Vacuum chamber
Sample
Feedback
Control software